Connector to the
transducer
IOs headers to
the FPGA IOs
High Voltage level
selector (75, 50, 25V)
Access to logic of
signal acquisition
Jumpers to FPGA
IOs, lowest one is
used to program FPGA
Green are headers
& connectors to
solder
Raspberry Pi 20x2 header

---
Connector to the
transducer
IOs headers to
the FPGA IOs
High Voltage level
selector (75, 50, 25V)
Access to logic of
signal acquisition
Jumpers to FPGA
IOs, lowest one is
used to program FPGA
Yellow are headers
& connectors to
solder
Raspberry Pi 20x2 header

---
Ice40
Flash
To program the FPGA,
place a jumper as indicated (in blue) on the « SPI FT if IN », which means you can program the FPGA flash with iceCube (Lattice) or iceprog (yosys)
Don’t forget to plug in using the USB beforehands =)
FTDI

---
PMOD
- 6 buffered inputs
- Logic Control access
- 3.3V and 5V
8 Mbit SRAM, 10ns, 512 k x 16, equivalent:
- 65 full lines of 120us at 64Msps
- 840 lines of 120us at 10Msps, 8 bits
High voltage pulser
- Pulser: MD1210 + TC6320
- High voltage at 25V, 50V, 75V
(selected using a jumber)
Analog to Digital Converter ADC10065
- 10-Bit 65 MSPS 3V A/D Converter
- Sampling at 64/(1+n) Msps
Time Gain Compensation (TGC)
- AD8331 controled by DAC
- 8 bit, 200us TGC programmable
3 x push controls button + LEDS
+ 1 reset button
FPGA
- Lattice iCE40HX4K
- TQFP 144 Package
Raspberry Pi 20x2 Header
8 Mb SPI Flash
for FPGA configuration
SMA dual placement for piezo connection
PMOD
- 8 IOs
- 4 Ios shared with RPi
Four jumpers
- 3 programmable
- 1 used for RPi / USB connection
USB connector (Power supply – needs 400mA, programmation)
FT2232H interface
Open source
Ultrasound acquisition and processing with open source hardware : un0rick.cc

---
TO BMP